BIX/CPR 230型

半身电脑心肺复苏模拟人

使用手册

Half-body Electronic CPR Manikin
User’s Manual
第七代CPR人体模型是根据美国心脏协会（AHA）2015年国际心肺复苏（CPR）和心血管护理（ECC）指南标准开发的新产品，是具有先进功能的国内领先产品， 外观美观，使用方便。The 7th generation  CPR manikin is a new product which is developed upon the requirement of Guide Standard of American Heart Association (AHA)2015International Cardiopulmonary-resuscitation (CPR) & Emergency Cardiovascular Care (ECC) and it is the domestically leading product with advanced functions, beautiful shape and convenient application. 
其功能特点是：模拟标准气道开放，外部胸部压迫，呼吸等主要功能及其操作程序分为：训练和试验。 此外，它在国内外客户中非常受欢迎。
The function characteristics of it are: simulating standard airway open, External Chest Compression, Respiration and other main functions and its operation procedures are divided into: training and test. Moreover, it owns great popularity among the clients both at home and abroad.
心肺复苏模拟人安装过程：

Operation Instruction of  
. Installation process of the manikin
先将模拟人从皮箱内取出，把模拟人平躺仰卧在操作台上，另将电脑显示器连接电源线，外接电源线从皮箱内取出，再与人体进行连接，将电脑显示器与220V电源接好，即完成连线过程
 Firstly, take the manikin out from leather suitcase and then lie it on the operating platform, and then connect the controller to the power cord. Finally, connect the cable between the controller and the artificial human body. Thus all the connection processes are finished. 

操作前功能设定及使用方法：

Function setting and application method before operation:  
完成连线过程后，打开电脑显示器后面的电源开关，此时显示器响起语音提示：“欢迎使用本公司产品，请选择工作方式”。
After finishing connecting wires, turn on the power switch at the back of controller, at this time, the display will release voice indication: “Welcome to use the manikin, please select working mode”. 
本产品有二种工作方式：

The product possesses two kinds of working modes: 

①训练：用户可以随意的进行人工呼吸和胸外按压操作，以熟练其技术
① training: user can carry out Respiration and External Chest Compression operation arbitrarily to get familiar with its technologies.
②考核：根据最新标准设定的考核模式。

②Test: carry out operation according to the evaluation mode set up upon the latest standard. 

完成工作模式选择后，语音提示：“请按开始按钮”。 此时可以执行操作。
  After finishing the working mode selection, the voice indication will be released: “Please press the button start”. Operation can be implemented at this time. 

Other keying functions: 其他按键功能：

1 复位键功能：既选定工作方式，按程序操作，操作不成功或其他原因需重新操作，请按一下复位键重新按先前设定好的方式操作，如需更改工作方式，请按住复位键5秒钟以上，即可重新选择工作方式。

Reset function: that means to select working mode and begin to operate according to the procedure. If it needs operation again duo to unsuccessful operation or other reasons, please press button "start" for more than 3 seconds to reset and operate again. 

操作过程中，必须要掌握规范动作及注意事项
Some specified actions and precautions to be grasped during the operation process 

胸外按压功能提示：首先找准胸部正确位置：胸骨下切向上两指胸骨正中部（胸口剑突向上两指处）为正确按压区，双手交叉叠在一起，手臂垂直于模拟人胸部按压区，进行胸外按压。若按压位置正确，按压强度正确（正确的按压深度为5-6厘米），人体按压条码灯的绿灯发光管（正确区域）显示，正确按压数码计数1次，若按压位置错误，将有语音提示：“按压位置错误”。若按压位置正确而按压强度错误，人体按压条码灯的黄灯发光管（不足）或红灯发光管（过大）显示，将有语音提示“按压不足”或“按压过大”等，按压错误数码计数1次，需纠正错误，再操作。

Function indication of chest compression: firstly, find the correct position on the chest which is in the middle of the nipple line two-fingers width in the upper of sternum undercut ,and this is the correct compressing area. Overlap and cross two hands, and the arms shall be vertical to the chest compressing area of the manikin  to carry out chest compression. If both of the compressing area and strength are correct (the correct compressing depth is 5-6 cm), the green light will display and the correct compression number will be counted once. While if the compressing area is not correct, there will be voice indication: “wrong compressing position”. If the compressing area is correct while the compressing strength is wrong, the yellow light led of human body compression bar code lamp (insufficient) or red light led (excessive) will display with voice indication “not enough” or “too much” etc. The compression wrong number will be counted once. It needs to correct the mistakes to operate again. 

2、人工呼吸功能提示：在进行口对口人工呼吸时，当操作者吹入的潮气量达到500-1000毫升，人体吹气条码灯的绿灯发光管（正确区域）显示，吹气正确数码计数1次。当操作者吹入的潮气量小于500或大于1000毫升，人体吹气条码灯的黄灯发光管（不足）或红灯发光管（过大）显示，吹气错误数码计数1次，并有语音提示：“吹气不足”或“吹气过大”等，需纠正错误后，再操作
1. Function indication of Respiration: while carrying out mouth-to-mouth  Respiration, if the quantity of the insufflates air reaches 500/600-1000 ml, the green light will display(correct area) and the air-blowing correct number will be counted once. When the quantity of the insufflates air is less than 500 or more than 1000 ml, the yellow light will display(insufficient) or red one display (excessive) and the wrong number of blowing will be counted once with voice indication: “not enough” or “too much” etc. It needs to correct the mistakes to operate again. 

四、操作方式：（特别提示：根据最新国际抢救修订标准，全面推行单人心肺复苏抢救标准步骤，按压频率采取最少100次/分。）

IV. Operation mode (special instructions: according to the latest revised standard of international emergency, the solo cardiopulmonary-resuscitation emergency standard procedure shall be implemented fully and the compressing frequency adopts 100 times/ minute.) 

1.训练：此项操作是让学员熟练掌握操作基本要领及各项步骤。当功能设定好，学员就可以进行人工呼吸或胸外按压。操作正确错误有各类功能数码显示及语音提示。

Training: this operation means to make the students get familiar with basic essentials and various procedures of operation. When the functions are set up, the students can carry out Respiration or chest compression. There will be various types of functional digital display and voice indication to show whether the operation is right or not. 

2. 考核：标准的考核标准。学员在经过训练操作，能熟练掌握了急救操作的基础上进行考试。学员必须按考试标准程序进行。

Test: the evaluation standard is established against the training staff. The students will take examination on the basis of grasping the emergency operation after operation training. There will be various types of functional digital display and voice indication to show whether the operation is right or not. 
单人考核按考试标准电脑程序操作的规范步骤：

Single evaluation will refer to the examination standard specification procedure operated by the computer steps: 

[image: image1.jpg]



步骤①——进行单人正确胸外按压30次（显示器上显示按压显示为30）

Step1----then carry out single person correct chest compression for 30 times (the compression display on the display will be 30) 

步骤②——再进行单人正确人工吹气2次（显示器上正确吹气显示为2）

Step2---carry out single person correct blowing for twice .
步骤③——连续进行正确胸外按压30次，正确人工呼吸2次（即30：2）

的五个循环（包括步骤①、②的一个循环在内）

  Step3---carry out correct chest compression successively for 30 times and 5 circulations (including one circulation in Step1 and 2) of twice correct artificial respiration (that is 30:2)
步骤④——最后显示器上正确按压显示为150，正确吹气显示为10。即告单人操作按程序操作成功，随之有语音提示：“操作成功”，自动奏响音乐，颈动脉连续搏动，心脏跳动，恢复心跳声音，瞳孔由原来的散大自动缩小，说明人被救活。

Procedure4---the correct compression on the display finally will be 150 and the correct air-blowing will display 10. That means the solo operation with procedure operation is successful and the person is saved. 

常见故障排除
Common Troubleshooting

现象一 ：开机后，设备不能启动或语音提示请检查模拟人连接
Phenomenon ①: the equipment fails to start up or the voice indication shows that “check the cable with the manikin and press'reset'” 

【检查一】：检查电源指示灯是否亮
[Inspection 1]: inspect whether the power indicator is on
【检查二】：检查模拟人连接线是否连接

 [Inspection 2]: inspect whether the cable of the manikin is connected

【检查三】：检查电源线插头是否插紧或脱落
[Inspection 3]: inspect whether the plug of power cord is plugged or detached 

【检查四】：检查语音旋扭开关是否打开

[Inspection 4]: inspect whether the voice button switch is turned on
【检查五】：检查保险管是否熔断
  [Inspection 5]: inspect whether the protective tube is fusing 

现象二：人工吹气后，电子显示器显示异常或语音提示异常

Phenomenon ②: the controller display or the voice indication is abnormal after blowing 

【检查一】：检查模拟人气道是否打开
[Inspection 1]: inspect whether the airway is open 

【检查二】：检查人工吹气方法是否正确
[Inspection 2]: inspect whether the method of blowing is correct

【检查三】：检查进气管是否脱落(取下脸皮后，即可在下颌处看见白色进气管)
[Inspection 3]: inspect whether the air inlet pipe is detached (after getting down the face skin, the white inlet pipe can be seen under the jaw)

【检查四】：检查肺气袋是否漏气(取下胸皮后，可在按压板下看见白色肺气袋)
 [Inspection 4]: inspect whether the lung bag leaks (after getting down the chest skin, the white lung bag can be seen under the compression plate) 

现象三：人工按压后，电子显示器显示异常或语音提示异常

Phenomenon ③: the controller display or the voice indication is abnormal after  compression 

【检查一】：检查按压板连接线或传感器连接线是否断裂

[Inspection 1]: inspect whether the connection wire of compression plate or the sensor breaks down

【检查二】：检查传感器内按压拉伸弹簧是否断裂

[Inspection 2]: inspect whether the compression extension spring in the sensor breaks down 

【检查三】：检查传感器内是否有灰尘

[Inspection 3]: inspect whether there is some dust within the sensor 
【检查四】：检查按压位置是否正确
[Inspection 4]: inspect whether the compression position is correct 
【检查五】：检查按压强度是否正确
[Inspection 5]: inspect whether the compressing strength is correct 

【检查六】：检查按压频率是否按频率要求
 [Inspection 6]: inspect whether the compressing frequency is in accordance with the frequency requirements 
现象四：成绩报告单无法打印或打印异常
【检查一】：检查打印机指示灯是否亮
【检查二】：检查打印面是否装反了（正常为光面操上）
【检查三】：检查打印纸前端是否露出打印仓外，自检走纸是否正常
【检查四】：打印纸是否用完或卡纸，如有请更换打印纸，打印纸光面朝上进行更换。排除故障后，重新启动。

Phenomenon ④: insufficient compression language indications  

[Inspection]: open the chest and lift off the screws in the middle part, then press the black sensor under it by hands to check whether it owns elasticity. If yes, the fault is probably caused by the dusts. Please open the black sensor and then brush the dust clearly with brush. The screws shall not be too tight or loosen when installing sensor. It needs to restart after shooting the trouble. 

现象五： 按压语言提示不足

【检查】：打开胸部，卸下中间螺丝，将下面黑色传感器用手按压一下，查看有没有弹性，若有则是因为灰尘导致故障，请拆开黑色传感器，将里面的灰尘用刷子轻轻刷干净即可。

安装传感器时螺丝不能过紧或过松。排除故障后，重新启动。
Phenomenon ⑤: no compression display 

[Inspection]: open the chest and lift off the screws in the middle part, then press the black sensor under it by hands to check whether it owns elasticity. If not, Please open the black sensor and change the spring in it. The screws shall not be too tight or loosen when installing sensor. It needs to restart after shooting the trouble. 

现象六： 按压没有显示

【检查】：打开胸部，卸下中间螺丝，将下面黑色传感器用手按压一下，查看有没有弹性，若没有则请拆开黑色传感器，更换弹簧即可。安装传感器时螺丝不能过紧或过松。排除故障后，重新启动。

Phenomenon ⑥: there is no voice indication between the connection of host and simulated man 

[Inspection]: inspect whether the power cord falls out or not well plugged, whether the power is normal. Inspect the sound volume at the back of host to check whether it opens or is loosened. It needs to restart after shooting the trouble.







